Title
Respiratory Motion Estimation –
Tests and Comparison of different Sensors
Inter-Disciplinary Project (Physics)
Student: Michael Riedel
Supervision: P. Böni1, N. Navab2, A. Martinez-Möller3, R. Bundschuh3
19 May, 2006
2 Chair for Computer Aided Medical Procedures & Augmented Reality | Department of Informatics | Technische Universität München
3 Department of Nuclear Medicine | Klinikum rechts der Isar | Technische Universität München
1 Institute for Experimental Physics E21 | Physics Department | Technische Universität München

---
Respiratory Gating
Context: Respiratory Gating in PET Imaging
  Introduction

---
Contents1
Respiratory Motion Estimation –
Overview
Introduction
Overview of PET Imaging
Tests of Respiration Sensors
Evaluation of Respiration Curves
Conclusion
  Overview of PET Imaging

---
Functional Imaging
Functional Imaging vs. Structural Imaging
PET Image
CT Image
Fused Image
Positron Emission Tomography (PET) is a functional imaging technique allowing quantification of tracer distribution
  Overview of PET Imaging

---
Positron Emmission
Principles of PET –
Why is it Suited for Functional Imaging?
PET is based on β+ decay:
Works with radionuclides and is independent of chemical processes
Radionuclides are built into tracers with various chemical and biological properties
Tracers applied to a patient are used to track molecular processes in vivo
  Overview of PET Imaging
"p" ~ rightarrow ~~~ "n" ~ + ` "e"^{"+"} ` + ` %nu_{"e"}
"X" lsup{"a"} lsub{"z"} ~ rightarrow ~ "Y" lsup{"a"} lsub{"z-1"} ` + ` "e"^{"+"} ` + ` %nu_{"e"}
"F" lsup{18} lsub{9} ~ rightarrow ~``` "O" lsup{18} lsub{8} ` + ` "e"^{"+"} ` + ` %nu_{"e"}
FDG
NH3
O2
H2O
Acetate

---
Scanner Geometry
PET Scanner Geometry
positron range
~ 2 mm
"e"^{"+"} ` + ` "e"^{"-"} ~ rightarrow ~ %gamma ` + ` %gamma
Positrons annihilate with nearby electrons:
511 keV γ rays are emitted in exactly opposite directions
γ rays originating from within
the patient can be detected
by the surrounding PET scanner
± 0.5°
angular deviation
  Overview of PET Imaging

---
Reconstruction
Reconstruction of PET Images (2D Case)
γ rays detected in temporal coincidence define a line of response (LOR)
Assumption: The emission of the γ rays occurred somewhere on the LOR
Tracer activity measured along each LOR can be organized in a sinogram
A sinogram is a Radon transform of the tracer distribution
Image reconstruction is done by inversion of the Radon transform, e.g. using:
Filtered Back-Projection
Statistical Iterative Reconstruction
Very similar to CT
image reconstruction!
  Overview of PET Imaging

---
Scintillators
Materials for PET Detector Rings –
Scintillators and Photomultiplier Tubes
Current PET scanners consist of a huge number of scintillator crystals
(e.g. Siemens Biograph PET/CT has 9216 crystals of 6.452 mm2 each)
Scintillators are based on Compton scatter between γ rays and electrons
Scintillators fluorescent light close to the visible spectrum
Emitted light can be detected by photomultiplier tubes
  Overview of PET Imaging

---
Degrading Effects
Degrading Effects and Limitations of Clinical PET Scanners
Only a small fraction of γ ray pairs reaches detectors, due to open geometry and attenuation
Many of those γ rays are scattered
(10% - 60%)
Scattered and random coincidences add noise
Scatter and attenuation are influenced by the patient's anatomy
  - Can be corrected by measuring an effective attenuation coefficient for each LOR in a separate transmission scan
Attenuation can also be measured indirectly by CT ⇒ PET/CT
Conclusion: The principles of PET are simple, but a huge variety of degrading effects has to be dealt with!
  Overview of PET Imaging

---
PET Summary
Summary of PET Imaging
Based on beta decay and annihilation:
Radionuclides are independent of
chemical processes, and can be built
into various PET tracers
A tracer applied to a patient beta decays
according to its distribution within the organism
γ rays are emitted in exactly opposite directions,
and can be detected by the PET scanner
Coincidence detections define a line of response – altogether yielding a sinogram representing a Radon transform of the tracer distribution
Reconstruction of 2D images slices, e.g. by filtered back projection
Very similar to CT image reconstruction!
Degrading effects like scatter and attenuation add complexity to PET!
"p" ` rightarrow ` "n" ` + ` "e"^{"+"} ` + ` %nu_{"e"}
"e"^{"+"} ` + ` "e"^{"-"} ` rightarrow ` %gamma ` + ` %gamma
  Overview of PET Imaging

---
Need for Gating
Why is Respiratory Gating Necessary?
The temporal resolution for detecting coincidences is limited
  - Higher tracer activity would cause problems due to dead-time
  - Higher tracer activity would cause multiple coincidences
  - Higher tracer activity might violate regulations
A lot of true coincidences are needed, to compensate for scatter (~107)
⇒ PET Scans generally take a long time!
(e.g. oncology scan with 18F-FDG: 3 minutes per bed position)
⇒ Respiratory motion is inevitable and needs to be corrected!!!
Characteristics of cardiac PET study used for evaluation of respiration sensors:
  - 300 - 500 Mbq of 13N-NH3 used as perfusion marker
  - PET acquisitions of 11 minutes each, also in list-mode
    (= raw data acquisition)
  - 2 PET scans per patient: rest and stress
  Overview of PET Imaging

---
Contents2
Respiratory Motion Estimation –
Overview
Introduction
Overview of PET Imaging
Tests of Respiration Sensors
Evaluation of Respiration Curves
Conclusion
The system is done.
We are just trying to get it to work.
Tests of Respiration Sensors

---
Respiration Sensors
Respiration Sensors
Anzai Belt (right):
Elastic belt measuring thorax/abdomen expansion
Based on a load cell and designed for respiratory gating
BioVet Temperature Sensor (left):
High precision thermometer with high temporal resolution
Intended for small animal research
Measures temperature of respired air
PMM Spirometer (not depicted):
Measures air-flow at the patients nose
Prototype system intended for respiratory gating
ART Stereo Infrared Camera (right):
Marker-based 3D tracking system
See next slides ...
Tests of Respiration Sensors

---
ART Stereo Infrared Camera
3D Tracking with the ART Stereo Infrared Camera
Based on the principles of stereo vision:
Depth perception by two different view-points
Cameras are mounted rigidly and are calibrated with respect to each other
Accuracy depends on actual setup, but is usually better than 1 mm; sampling frequency is 60 Hz
Uses retro-reflexive markers lighted by infrared flashes
Single markers only reveal location (3DOF)
Rigid bodies of markers reveal location and orientation (6DOF)
Tracking is performed transparently
Tests of Respiration Sensors

---
Clinical ART Setup
Clinical Setup of the ART Camera
Intended for measurement of respiration and body motion
5 markers (3DOF) fixed to abdomen and thorax
Camera is placed at the head-end of the PET/CT
Line of sight needs to remain free – problematic due to limited size of patient port
Mobile setup: needed to be remounted for each patient!
Tests of Respiration Sensors

---
ART Calibration
Coordinate System Calibration –
Registering ART Data with PET Images
ART tracking data are in arbitrary “room” coordinate system
Tracking data need to be registered with respect to the PET/CT scanner
Orientation was important – absolute location was not necessary, just relative
A tracked 6DOF reference body (B) was fixed permanently to the PET/CT scanner
The calibration matrix (C) transforming from PET/CT to reference body coordinates was obtained by calibration measurements
Tracking data were transformed accordingly:
{vec p}' ~ = ~ {bold "C"}^{"-1"} {bold "B"}^{"-1"} {vec p}
Tests of Respiration Sensors

---
ART Processing
From Tracked Markers to Respiration Curves
Tests of Respiration Sensors

---
Contents3
Respiratory Motion Estimation –
Overview
Introduction
Overview of PET Imaging
Tests of Respiration Sensors
Evaluation of Respiration Curves
Conclusion
98% of all statistics are made up.
Evaluation of Respiration Curves

---
Respiration Curves
Respiration Curves Reflecting Both
Regular Respiration and Breathing Artifacts
Evaluation of Respiration Curves

---
Gating Artifacts 1
Impact of Differing Signals on Respiratory Gating
Gates computed by detecting peaks and equally dividing the resulting intervals:
Evaluation of Respiration Curves

---
Gating Artifacts 2
Impact of Differing Signals on Respiratory Gating
Gates computed by detecting peaks and equally dividing the resulting intervals:
Evaluation of Respiration Curves

---
Statistical Comparison
Statistical Comparison of the Respiration Sensors
Correlation(R,S) ~ = ~ {"E"((R - %mu_{R})(S - %mu_{S}))} over {%sigma_{R} %sigma_{S}}
GatedCorrespondence(R,S) ~ = ~ {1 over n} sum from{i=1} to{n} left lbrace {stack{1 ` "if" gate(R_{i}) = gate(S_{i}) # 0 ` "otherwise "}} right rbrace
Caution: Poor underlying data!
High similarity under clinical conditions
Tends to get worse under stress
Reasonable similarity for all sensors
PMM is impaired by timing problems
Evaluation of Respiration Curves

---
Gated Images
Example of Gated PET Images
ART respiration curves were divided into 4 gates
Gating was applied to 11 minutes of list-mode PET data
Full expiration (Gate 3)
Full inspiration (Gate 1)
Images show the myocardium surrounding the left ventricle of the heart
Motion of the heart can be observed (here around 4 mm for the inferior wall)
Evaluation of Respiration Curves

---
Contents4
Respiratory Motion Estimation –
Overview
Introduction
Overview of PET Imaging
Tests of Respiration Sensors
Evaluation of Respiration Curves
Conclusion
Almost there ...
Conclusion

---
Respiration Sensor Requirements
Which Respiration Sensor to Choose?
All tested sensors basically work
Choice should be based on quality requirements like usability, invasiveness, reliability, and costs
  - Anzai belt fulfills that best!
  - BioVet thermometer and PMM spirometer have reliability problems
  - ART camera is comparatively cumbersome
BUT:
ART camera can be used for
quantitative measurements of
respiration and body motion...
Conclusion

---
Body Motion Estimation
Motion Estimation Beyond Respiratory Gating
ART allows quantification of respiratory motion on the patient's surface
  - Same order of magnitude as the motion of the heart seen in gated PET images (around 5 – 12 mm in average)
Low-pass filtering ART data, yields body motion not related to the respiration cycle
  - Patients typically do another 5 – 10 mm of irregular body motion
  - Body motion cannot be corrected
    by respiratory gating!
  - Interesting finding:
    Adenosine injection for PET scan
    under stress can increase body motion
Conclusion

---
Outcome and Future Work
What is done? What is Left?
Outcome of this IDP:
  - Applicability of four respiration sensors was proved
  - Clinical setup and software components for the ART camera as a respiration sensor were developed:
    The ART camera can now be used in clinical studies
    with real patients!
  - Software tools for evaluation and comparison of respiration curves and of body motion were created
Future Work:
  - Measure respiratory motion of more patients! (Just 4 by now!)
  - Evaluate recombination of gated PET images by applying non-rigid transformations
  - Examine body motion of different patients more thoroughly:
    Is translation, rotation, scaling or skew significant?
  - Try to correct general body motion, if necessary
Conclusion